ΔΙΑΓΩΝΙΣΜΟΣ ΓΙΑ

Προμήθεια, Εγκατάσταση και Συντήρηση

Μηχανήματος: Heat Flow Calorimeter

(HFC)

Aριθμός  Διαγωνισμού:  13.25.05.2013.18

ΤΟΜΟΣ Γ: ΤΕΧΝΙΚΕΣ ΠΡΟΔΙΑΓΡΑΦΕΣ

Τόμος Γ_Προμήθειες_Ιανουάριος_2011

## ΕΝΤΥΠΟ 14 41

# ΠΙΝΑΚΑΣ ΠΡΟΣΦΟΡΑΣ ΚΑΙ ΣΥΜΜΟΡΦΩΣΗΣ ΜΕ ΤΙΣ ΤΕΧΝΙΚΕΣ ΠΡΟΔΙΑΓΡΑΦΕΣ

### Ζητούμενο προϊόν (1): ___Heat Flow Calorimeter (HFC) suitable for analyzing Nitrocellulose based propellants in

### accordance with the International NATO Stanag STANAG 4582

| Α/Α | Χαρακτηριστικά - Προδιαγραφές | Απαίτηση | Απάντηση - Προσφορά | Παραπομπή |
| --- | --- | --- | --- | --- |
|  | ΓΕΝΙΚΑ ΧΑΡΑΚΤΗΡΙΣΤΙΚΑ |
|  | Αριθμός μονάδων | 1 |  |  |
|  | Κατασκευαστής – Μοντέλο |  |  |  |
|  | ΕΙΔΙΚΑ ΧΑΡΑΚΤΗΡΙΣΤΙΚΑ |
|  | Heat Flow Calorimeter - measuring range and maximum error during the total measuring time (typically several days). | Range10-500μW/g Error < 2% |  |  |
|  | The instrument must be capable of housing sample vessels with a minimum volume of 2 cm3 that can be hermetically sealed. | Yes |  |  |
|  | It must be proved that the vessel material does not react with propellants and their decomposition products. | Yes |  |  |
|  | An appropriate data collection and calculation system including integration is required. | Yes |  |  |
|  | Capacity of the system for multiple HFC measurement to be undertaken simultaneously(minimum 24) | Yes |  |  |
|  | THERMOSTAT |  |  |  |
| 2.6 | Temperature range | 15oC-150oC |  |  |
| 2.7 | Temperature Accuracy | <+/- 0.1K |  |  |
| 2.8 | Long Term Stability | <+/- 100μK/24 hr |  |  |
| 2.9 | Short Term Stability | < +/-10μK(peak to peak) |  |  |
| 2.10 | Scanning Rate | <+/-2K/h(between 20-150 oC |  |  |
| 2.11 | Heating Rate | 15K/h at 15 oC-2K/h at 150 oC |  |  |
| 2.12 | Cooling Rate | 15K/h at 150 oC-1.5K/h at 15 oC |  |  |
| 2.13 | Thermostat must be able to measure the samples specified in section 2.5 simultaneously and be expandable to measure up to a maximum of 48 samples | Yes |  |  |
|  | CALORIMETERS SHOULD INCLUDE THE FOLLOWING MODULES | Yes |  |  |
| 2.14 | Microcalorimeter: for isothermal or temperature scanning measurements on 20ml or 4ml sample sizes | Yes |  |  |
| 2.15 | Mini/multicalorimeter: An array of minicalorimeters, each providing microcalorimeter-level sensitivity | Yes |  |  |
|  | MICROCALORIMETER SPECIFICATIONS |  |  |  |
| 2.16 | Dynamic Range: | ± 30 milliwatts |  |  |
| 2.17 | Short Term Noise: | < ±100 nanowatts |  |  |
| 2.18 | Baseline Drift: | < 200 nanowatts/24 hours |  |  |
| 2.19 | Accuracy: | < 2% |  |  |
| 2.20 | Precision | ± 100 nanowatts |  |  |
| 2.21 | The calorimeter must be a true isothermal system, not adiabatic. As such, the temperature of the sample should not change more than ±0.0001°C. | Yes |  |  |
| 2.22 | The calorimeter must operate in heat flow (heat flux) or dynamically corrected mode of detection. | Yes |  |  |
| 2.23 | The heat flow response of the calorimeter must be calibrated using a fixed calibration heater. | Yes |  |  |
| 2.24 | The calorimeter must accommodate removable sample ampoules, up to 20 mL in volume. | Yes |  |  |
| 2.25 | The calorimeter must accommodate a wide variety of ampoules, including the following: -Closed ampoules, made from either glass, stainless steel or hastelloy and employing either heat seals, screw caps or circlip seals -Perfusion ampoule, for introduction of a gas or a liquid to the sample during the experiment | Yes |  |  |
|  | 4 ml Mini/Multicalorimeter specifications |  |  |  |
| 2.26 | Microcalorimeter system must accommodate multiple simultaneous differential measurements. | Yes |  |  |
| 2.27 | Dynamic Range | ±25 milliwatts |  |  |
| 2.28 | Short Term Noise | : <±100 nanowatts |  |  |
| 2.29 | Baseline Drift | < 200 nanowatts/24 hours |  |  |
| 2.30 | Accuracy | < 5% |  |  |
| 2.31 | Precision | ±200 nanowatts |  |  |
| 2.32 | The calorimeter must be a true isothermal system, not adiabatic. As such, the temperature of the sample should not change more than ±0.0001°C. | Yes |  |  |
| 2.33 | The calorimeter must operate in heat flow (heat flux) mode of detection. | Yes |  |  |
| 2.34 | The heat flow response of the calorimeter must be calibrated using a fixed calibration heater. | Yes |  |  |
| 2.35 | The calorimeter must accommodate removable sample ampoules, up to 5 mL in volume. | Yes |  |  |
| 2.36 | The calorimeter must accommodate a wide variety of ampoules, including the following: -Closed ampoules, made from either glass, stainless steel or hastelloy and employing either heat seals, screw caps or circlip seals | Yes |  |  |
| 3 | SOFTWARE AND PC |  |  |  |
| 3.1 | Appropriate software must be included, able to control the system and analyze measurements according to STANAG 4582 International norm | Yes |  |  |
| 3.2 | Data processor operating system for full system control of the hardware with double hard disk, data acquisition, storage, retrieval, processing, reprocessing manipulation and others. The PC should be with Microsoft Windows 7 or XP with a recent operating system. Processor (CPU) i7,4GB RAM minimum supporting RAIDO 2X 1TB, double hard disk (main and mirror disk), DVD/CDROM, Multiburner DVD 20”colour wide TFT flat screen 3ms 3D Graphic card with minimum 512MB RAM Network compatible Switch Ethernet 10/100, 8-port Keyboard and optical mouse Colour Inkjet printer, high resolution, low colour printing cost. Automatic Back-up system. | Yes |  |  |
| 4 | Accessories |  |  |  |
| 4.1 | Start Up Kit 4 mL, with 50 Eyelets | YES |  |  |
| 4.2 | 500 Transparent glass, disposable chemically inert ampoules 4 mL suitable for stability and compatibility measurements . | YES |  |  |
| 4.3 | 1000 Alum. Caps, Sealing Discs, Glass Ampoules 3,4 mL | YES |  |  |
| 4.4 | 5 Gas sampling Bags | YES |  |  |
| 4.5 | OIL MOBILTHERM , 27L | YES | 1 |  |
| 4.6 | Digital Block Heater , suitable to be used for pre-heating samples and ampoules prior to measurement or to keep samples exposed to a defined temperature in between measurements. The block heater should be configured with two of 4 mL block inserts Each block insert must hold up to 24 ampoules . All ampoules which fit into the 4 mL Multicalorimeter must fit into these block inserts. Temperature stability of 0.1 °C Temperature range +5 °C from ambient up to 120 °C. | YES |  |  |
| 5 | UPS | Yes |  |  |
| 5.1 | UPS system to ensure the non-interrupted operation of the instrument | Yes |  |  |
| 6 | General Requirements |  |  |  |
| 6.1 | Prospective supplier is responsible for system installation and start up of test. | Yes |  |  |
| 6.2 | All necessary accessories needed to install and start up of system should be included. | Yes |  |  |
| 6.3 | Operational training for at least 2 users must be provided, for at least 5 days. On-site Service Training to be provide by the supplier. | Yes |  |  |
| 6.4 | Suppliers should provide with offer, manufacturers authorization, authorizing them to sell and service the offered system in Cyprus. | Yes |  |  |
| 6.5 | Suppliers should provide with offer a list of at least one local but factory trained service engineer, capable to service and support the offered system. Training certificate must be provided with the offer. | Yes |  |  |
| 6.6 | 1 Service and 2 set Operators Manuals to be supplied with the instrument. | Yes |  |  |
| 6.7 | The tender to include undertaking of the removal and relocation and installation of all parts of the instrument, should this become necessary in the future | Yes |  |  |
| 7 | WARRANTY and SERVICE |  |  |  |
| 7.1 | 12 months warranty minimum | Yes |  |  |

ΣΗΜΕΙΩΣΕΙΣ:

  - Είναι υποχρεωτική η απάντηση σε όλα τα σημεία του Πίνακα. Σε περίπτωση που δεν έχει απαντηθεί οποιοδήποτε σημείο, η απάντηση θεωρείται αρνητική.

  - Στη Στήλη «ΧΑΡΑΚΤΗΡΙΣΤΙΚΑ-ΠΡΟΔΙΑΓΡΑΦΕΣ» ή «ΑΠΑΙΤΗΣΕΙΣ-ΠΡΟΔΙΑΓΡΑΦΕΣ», περιγράφονται αναλυτικά οι αντίστοιχοι τεχνικοί όροι, υποχρεώσεις ή επεξηγήσεις για τα οποία θα πρέπει να δοθούν αντίστοιχες απαντήσεις.

  - Αν στη στήλη «ΑΠΑΙΤΗΣΗ» έχει συμπληρωθεί η λέξη «ΝΑΙ», που σημαίνει ότι η αντίστοιχη προδιαγραφή είναι υποχρεωτική για τον υποψήφιο Ανάδοχο ή ένας αριθμός που σημαίνει υποχρεωτικό αριθμητικό μέγεθος της προδιαγραφής και απαιτεί συμμόρφωση, θεωρούνται ως απαράβατοι όροι και οι Προσφορές που δεν καλύπτουν τέτοιες απαιτήσεις απορρίπτονται ως απαράδεκτες.

  - Στη στήλη «ΑΠΑΝΤΗΣΗ» σημειώνεται η απάντηση του Προσφέροντα που είτε είναι το συγκεκριμένο προϊόν που προσφέρει είτε έχει τη μορφή ΝΑΙ/ΟΧΙ εάν η αντίστοιχη προδιαγραφή πληρούται ή όχι από την Προσφορά ή ένα αριθμητικό μέγεθος που δηλώνει την ποσότητα του αντίστοιχου χαρακτηριστικού στην Προσφορά.

  - Στη στήλη «ΠΑΡΑΠΟΜΠΗ» θα καταγραφεί η σαφής παραπομπή σε Παράρτημα της Τεχνικής Προσφοράς το οποίο θα περιλαμβάνει αριθμημένα Τεχνικά Φυλλάδια κατασκευαστών, ή αναλυτικές τεχνικές περιγραφές των προσφερόμενων προϊόντων, που κατά την κρίση του Προσφέροντα τεκμηριώνουν τα δηλούμενα στον Πίνακα Στοιχεία. Οι παραπομπές πρέπει να είναι συγκεκριμένες (π.χ. Τεχνικό Φυλλάδιο 3, Σελ. 4 Παράγραφος 4, κ.λ.π.), ενώ αντίστοιχα στο τεχνικό φυλλάδιο ή την αναφορά θα υπογραμμίζεται το σημείο τεκμηρίωσης και θα σημειώνεται η αντίστοιχη παράγραφος του Πίνακα στην οποία καταγράφεται το ζητούμενο χαρακτηριστικό.

  - Αφορά την παράγραφο 6.6 – Ο προμηθευτής θα πρέπει κοινοποιεί όποιες πληροφορίες του κοινοποιούνται από τον κατασκευαστή σε θέματα λειτουργίας ή συντήρησης οι οποίες δεν αναφέρονται στα εγχειρίδια που έχει προμηθεύσει.

ΣΗΜΕΙΩΣΕΙΣ ΓΙΑ ΤΗΝ ΑΝΑΘΕΤΟΥΣΑ ΑΡΧΗ Η στήλη «Χαρακτηριστικά - Προδιαγραφές» ή «Απαιτήσεις – Προδιαγραφές» συμπληρώνεται κατά περίπτωση και περιλαμβάνει συγκεντρωμένες τις απαιτήσεις, τόσο σε ότι αφορά στις ποσότητες και το είδος των ζητούμενων προϊόντων ή υπηρεσιών όσο και σε ότι αφορά τα ειδικά χαρακτηριστικά ή απαιτήσεις ή προδιαγραφές που αυτά/ές πρέπει να πληρούν. Στη στήλη «Απαίτηση» συμπληρώνεται είτε η λέξη «ΝΑΙ» που σημαίνει ότι η αντίστοιχη προδιαγραφή είναι υποχρεωτική για τον υποψήφιο Ανάδοχο ή ένας αριθμός που σημαίνει υποχρεωτικό αριθμητικό μέγεθος της προδιαγραφής και απαιτεί απόλυτη συμμόρφωση. Είναι δυνατό η στήλη «Απαίτηση» να περιλαμβάνει αριθμό με την ένδειξη «>» ή «<» ή «≥» ή «≤» ενώ επιτρέπεται και να μην είναι συμπληρωμένη για κάποιο χαρακτηριστικό ή προδιαγραφή, σε περίπτωση που αυτό το χαρακτηριστικό ή η προδιαγραφή δεν είναι υποχρεωτικό αλλά προαιρετικό και αξιολογούμενο. Σε περίπτωση που ως κριτήριο ανάθεσης έχει επιλεγεί η χαμηλότερη τιμή, στις τεχνικές προδιαγραφές θα πρέπει να ζητούνται μόνο τα απαραίτητα (υποχρεωτικά) χαρακτηριστικά – προδιαγραφές του προϊόντος ή οι απαραίτητες απαιτήσεις – προδιαγραφές των υπηρεσιών.